## Slide 1: Metal-conscious Transformer Enhanced CBCT Projection Inpainting
Reporter:
Yangkong Wang
Advisor :
M. Sc. Fuxin Fan, Prof. Dr.-Ing. habil. Andreas Maier

## Slide 2
Theoretical Background
2
Results and Discussion
4
Motivation
1
Methods
3
Summary
5

## Slide 3
Motivation
01
Part

## Slide 4
Metal artifacts
Fig.1: The existence of metal causes heavy streak artifacts in the reconstruction image. [1]
The photons of X-rays are scattered due to the interaction of the photons with the atoms of various body tissues.
When X-rays pass through dense metallic materials, the low-energy photons decay and the proportion of high-energy photons increases.
The increased absorption of photons by metals leads to a significant attenuation of X-rays.
Partial Volume Effect
Beam Hardening Effect
Scattering Effect

## Slide 5
MAR methods comparison
MAR
Metal artifact reduction algorithms are used to improve CT/CBCT image quality in patients with metal implants.
Sinogram/Projection Domain
Linear Interpolation (LI) [2]
Normalized MAR (NMAR) [3]
Frequency split MAR (FSNMA) [4]
Mask Pyramid Network (MPN) [5]
Gated convolutions MPN [6]
Dual Domain
Dual Domain Network for CT Metal Artifact Reduction (DuDoNet) [10]
Dual-Domain Adaptive-Scaling Non-local Network for CT Metal Artifact Reduction (DAN-Net) [11]
PriorNet-SinoNet [12]
Image Domain
DestreakNet [7]
Residual Learning Method based on Convolutional Neural Network (RL-ARCNN) [8]
Deep Interpretable Convolutional Dictionary Network for Metal Artifact Reduction in CT Images (DICDNet) [9]

## Slide 6
Overall projection domain-based MAR Pipeline
CBCT image MAR:
 Projection domain
 Image domain (truncation artifacts  limited detector range)
 Dual domain (huge computational complexity)
Metal-Corrupted Projection
2D Metal Mask Projection
1）
Inpainting Method
Inpainted Projection
2）
Metal-Free Reconstruction
3D Metal Mask
Final MAR Reconstruction
Metal Insertion
3）

## Slide 7
Overall projection domain-based MAR Pipeline
Metal-Corrupted Projection
Metal-Corrupted Reconstruction
Final MAR Reconstruction
2D Metal Mask
Segmentation Network
2）
Inpainting Network
Inpainted Projection
3）
Metal-Free Reconstruction
4）Metal Insertion
1） Threshold-based 3D Segmentation
3D Metal Mask
CBCT image MAR:
 Projection domain
 Image domain (truncation artifacts  limited detector range)
 Dual domain (huge computational complexity)

## Slide 8
Theoretical Background
02
Part

## Slide 9
Linear Interpolation
Input Projection
Inpainted Projection
Algorithm
Metal-corrupted Projection
Metal Projection

## Slide 10
MPN - Mask Pyramid Network [5]
Mask Pyramid Network (MPN)
CNN Block
Average Pooling
Concatenate

## Slide 11
Gated convolutional MPN [6]
0.04 | 0.05 | 0.04
0.05 | 0.07 | 0.05
0.04 | 0.05 | 0.04
0.04 | 0.05 | 0.04
0.05 | 0.07 | 0.05
0.04 | 0.05 | 0.04
Conv (k3,s1,p1)
BN
ReLU
Standard CNN Block
Conv (k3,s1,p1)
BN
ReLU
Gated CNN Block
Conv (k3,s1,p1)
Sigmoid
`
Gate
0.04 | 0.05 | 0.04
0.05 | 0.07 | 0.05
0.04 | 0.05 | 0.04
`
0.04 | 0.05 | 0.04
0.05 | 0.07 | 0.05
0.04 | 0.05 | 0.04

## Slide 12
Hybrid ViT-CNN network refer to SETR (Semantic Segmentation Transformer) [13]
Embedded Patches
Norm
Multi-Head Attention
Norm
MLP
×12
Vision Transformer Block
Pixel-wise Multiplication
ViT Block
CNN Block
Max Pooling
Positional Embedding
Reshape
Image Embedding
Concatenate
Detail description:
Image size = 512×512,
Patch size = 16×16,
Embedding dimension = 512,
Depths = 12,
Number of head = 8.

## Slide 13
Shifted window Vision Transformer (Swin) [14]
A local window to perform self-attention
A patch
(a) Vision Transformer
Patch Merging
Patch Merging
(b) Swin Transformer
Shifted window approach
Split
&
Concatenate
Patch merging approach
Flatten
&
LN
&
Linear

## Slide 14
Methods
03
Part

## Slide 15
Baseline network – The hybrid Swin-CNN network for inpainting task
Detail description:
Image size = 512×512,
Patch size = 4×4,
Embedding dimension = 128,
Depths = [2, 4, 4],
Number of heads = [4, 8, 8],
Window size = 8.
Reshape
Patch Merging
Image Embedding
Concatenate
Swin Block
CNN Block
Max Pooling
Positional Embedding
Pixel-wise Multiplication

## Slide 16
Baseline network with metal-conscious embedding
Metal Embedding
Reshape
Patch Merging
Image Embedding
Concatenate
Swin Block
CNN Block
Max Pooling
Metal Block
Positional Embedding
Pixel-wise Multiplication

## Slide 17
Baseline
Metal-conscious self-embedding method (SE)
Reshape
Patch Merging
Image Embedding
Concatenate
Swin Block
CNN Block
Max Pooling
Metal Block
Positional Embedding
Pixel-wise Multiplication

## Slide 18
Baseline
Metal-conscious neighborhood-embedding method (NE)
Reshape
Patch Merging
Image Embedding
Concatenate
Attention
Swin Block
CNN Block
Max Pooling
Metal Block
Positional Embedding
Input Embedding
Pixel-wise Multiplication

## Slide 19
An example of a metal attention matrix for the neighborhood embedding at the window size of 4.
…
…
…
…
…
…
… …
…
16
0 0 0 ··· 0 0
1 1 1 ··· 1 1
1 1 1 ··· 1 1
0 0 0 ··· 0 0
1 1 1 ··· 1 1
0 0 0 ··· 0 0
··· ···
0 0 0 ··· 0 0
16
0 0 0 ··· 0 0 0
…
…
0 0 0 ··· 0 0 0
…
0 0 0 ··· 0 0 0
… …
0 0 0 ··· 0 0 0
16
4
4
4
4

## Slide 20
Conditional Generative Adversarial Network (cGAN) [15]
Content loss:
Adversarial loss:
Total loss:
G
D
fake / real
N
CNN Block
Average Pooling
Pixel-wise Multiplication

## Slide 21
Results and Discussion
04
Part

## Slide 22
Experimental Parameters
Parameters for the mobile C-arm system (Simulation dataset) | 
Parameter | Value
Scan angular range | 360°
Incremental angular step | 6°
Source-to-detector distance | 1164mm
Source-to-isocentor distance | 622mm
Detector size | 512×512
Detector pixel size | 0.580mm×0.580mm
Hyperparameters for network training | 
Parameter | Value
λ | 100
Batch size | 10
Learning rate | 
Scheduler | 
Optimizer | 
Epochs | 100
Hyperparameters for network training | 
Parameter | Value
λ | 100
Batch size | 10
Learning rate | 
Scheduler | 
Optimizer | 
Epochs | 100

## Slide 23
Quantitative inpainting results (Simulation dataset : 3000projections = 2700Training + 300Validation)
 | Projection |  |  | Reconstruction |  | 
Method | MAE | PSNR | SSIM | MAE (HU) | PSNR | SSIM
Interpolation | 0.135 | 37.204 | 0.966 | 25.403 | 35.501 | 0.928
MPN | 0.111 | 39.606 | 0.970 | 25.089 | 36.394 | 0.939
MPN (gated) | 0.096 | 40.685 | 0.974 | 25.054 | 37.466 | 0.941
ViT-CNN | 0.295 | 31.845 | 0.932 | 47.156 | 30.617 | 0.850
Swin-CNN (Baseline) | 0.098 | 40.529 | 0.973 | 23.187 | 37.321 | 0.943
MPN
  Effectively handle complex metal implants
  Limited receptive field
Gated convolutions MPN
  Dynamic information extraction (gating mechanism)
  Limited receptive field
Hybrid ViT-CNN
  Global receptive field (self-attention)
  Not conducive to regression tasks (larger patch)
Hybrid Swin-CNN
  Local information (smaller patch, window self-attention)
  Global receptive field (patch merge, shifted window self-attention)
Linear Interpolation
  Simple algorithm
  Insufficient assumption
Conclusions:

## Slide 24
Quantitative inpainting results (Simulation dataset : 3000projections = 2700Training + 300Validation)
 | Projection |  |  | Reconstruction |  | 
Method | MAE | PSNR | SSIM | MAE (HU) | PSNR | SSIM
Baseline | 0.098 | 40.529 | 0.973 | 23.187 | 37.321 | 0.943
SE | 0.088 | 41.323 | 0.976 | 19.081 | 38.227 | 0.957
NE W4 | 0.092 | 40.985 | 0.975 | 20.201 | 38.321 | 0.955
NE W8 | 0.087 | 41.381 | 0.977 | 18.625 | 38.780 | 0.958
NE W16 | 0.079 | 42.346 | 0.978 | 18.021 | 38.921 | 0.960
NE W32 | 0.088 | 41.237 | 0.976 | 20.910 | 38.333 | 0.953
0.080
0.082
0.084
0.086
0.088
0.090
0.092
0.092
0.088
0.087
0.079
Window size
MAE
4
8
16
32
Metal-conscious embedding methods benefit baseline network.
In the case of NE, the choice of window size matters.
  Small window does not consider enough surrounding patches.
  Large window, irrelevant patches will account for the embedding of the unknown region.
Conclusions:

## Slide 25
Qualitative Inpainting results (Simulation dataset : 3000projections = 2700Training + 300Validation)
MPN (Gated CNN)
Hybrid ViT-CNN
MPN
Hybrid Swin-CNN
Interpolation
Label
Input
Metal Mask

## Slide 26
Qualitative inpainting results (Simulation dataset : 3000projections = 2700Training + 300Validation)
Hybrid Swin-CNN NE (W16)
Hybrid Swin-CNN NE (W8)
Hybrid Swin-CNN NE (W4)
Hybrid Swin-CNN NE (W32)
Hybrid Swin-CNN SE
Label
Input
Metal Mask

## Slide 27
Applying SE/NE W16 to the different stages of the baseline network encoder
 |  | SE | SE | SE | 
Stage 1 |  |  |  |  | 
Stage 2 |  |  |  |  | 
Stage 3 |  |  |  |  | 
Projection | MAE | 0.098 | 0.090 | 0.093 | 0.088
 | PSNR | 40.529 | 41.207 | 41.025 | 41.323
 | SSIM | 0.973 | 0.976 | 0.976 | 0.976
Reconstruction | MAE (HU) | 23.187 | 20.672 | 19.084 | 19.081
 | PSNR | 37.321 | 37.871 | 38.200 | 38.227
 | SSIM | 0.943 | 0.953 | 0.957 | 0.957
 |  | NE W16 | NE W16 | NE | 
Stage 1 |  |  |  |  | 
Stage 2 |  |  |  |  | 
Stage 3 |  |  |  |  | 
Projection | MAE | 0.098 | 0.089 | 0.083 | 0.079
 | PSNR | 40.529 | 41.314 | 41.821 | 42.346
 | SSIM | 0.973 | 0.976 | 0.978 | 0.978
Reconstruction | MAE (HU) | 23.187 | 18.840 | 18.602 | 18.021
 | PSNR | 37.321 | 38.518 | 38.903 | 38.921
 | SSIM | 0.943 | 0.958 | 0.960 | 0.960

## Slide 28
Combining SE with NE W16 using different ways
 |  | Projection | SE |  | Reconstruction | NE | 
Embedding |  | MAE | PSNR | SSIM | MAE (HU) | PSNR | SSIM
SE |  | 0.088 | 41.323 | 0.976 | 19.081 | 38.227 | 0.957
NE W16 |  | 0.079 | 42.346 | 0.978 | 18.021 | 38.921 | 0.960
SE + NE W16 |  | 0.097 | 40.608 | 0.975 | 22.622 | 37.485 | 0.946
 |  | 0.091 | 41.092 | 0.975 | 19.172 | 38.205 | 0.957
 |  | 0.085 | 41.662 | 0.977 | 19.047 | 38.648 | 0.957
 |  | Projection | SE |  | Reconstruction | NE | 
Embedding |  | MAE | PSNR | SSIM | MAE (HU) | PSNR | SSIM
SE |  | 0.088 | 41.323 | 0.976 | 19.081 | 38.227 | 0.957
NE W16 |  | 0.079 | 42.346 | 0.978 | 18.021 | 38.921 | 0.960
SE + NE W16 |  | 0.097 | 40.608 | 0.975 | 22.622 | 37.485 | 0.946
 |  | 0.091 | 41.092 | 0.975 | 19.172 | 38.205 | 0.957
 |  | 0.085 | 41.662 | 0.977 | 19.047 | 38.648 | 0.957
Metal-conscious self-embedding (SE)
Metal-conscious neighborhood-embedding (NE)

## Slide 29
Quantitative inpainting results (Cadaver dataset: 4000projections = 3600Training + 400Validation)
 | Projection |  |  | Reconstruction |  | 
Method | MAE | PSNR | SSIM | MAE (HU) | PSNR | SSIM
Interpolation | 0.099 | 42.082 | 0.984 | 20.759 | 37.426 | 0.949
MPN | 0.085 | 44.196 | 0.987 | 15.784 | 40.633 | 0.965
Gated MPN | 0.080 | 44.469 | 0.988 | 15.380 | 41.037 | 0.966
ViT-CNN | 0.254 | 35.605 | 0.974 | 31.243 | 34.539 | 0.903
Swin-CNN | 0.080 | 44.471 | 0.987 | 14.615 | 41.381 | 0.969
SE | 0.074 | 45.153 | 0.988 | 13.928 | 41.721 | 0.971
NE W16 | 0.069 | 45.471 | 0.988 | 13.755 | 41.841 | 0.972

## Slide 30
Summary
05
Part

## Slide 31
Summary
Linear Interpolation:
  Simple algorithm (speed and simplicity)
  Insufficient assumption (no global information)
Mask Pyramid Network (MPN):
  Ability to effectively handle complex metal implants (MPN)
  Limited receptive field
Gated convolutions MPN:
  Ability to effectively handle complex metal implants (MPN)
  Dynamic information extraction within convolutional receptive field (gating mechanism)
  Limited receptive field
Hybrid ViT-CNN:
  Large receptive field (self-attention approach)
  Not conducive to regression tasks (larger patch size)
Hybrid Swin-CNN:
  Local and global information (hierarchical structure, image split and merge)
  Shorter sequence, suitable for high-resolution downstream tasks (self-attention within each window, smaller patch size)
  Different region connections (shifted window self-attention)
Metal-conscious self-embedding method (SE):
  Providing a priori metal information, such as metal location, shape and size
  No connection between the metal region and the tissue region
Metal-conscious neighborhood-embedding method (NE):
  Reinforcing the connection between the metal and tissue regions
  Reasonably generated content (window design)

## Slide 32
Future work - Using dynamic window instead of fixed window
Future work :
unsuitable window
suitable window

## Slide 33
References
[1] Tristan M Gottschalk, Andreas Maier, Florian Kordon, and Björn W Kreher. View-consistent metal segmentation in the projection domain for metal artifact reduction in cbct–an investigation of potential improvement. arXiv preprint arXiv:2112.02101, 2021.
[2] Willi A Kalender, Robert Hebel, and Johannes Ebersberger. Reduction of ct artifacts caused by metallic implants. Radiology, 164(2):576–577, 1987.
[3] Esther Meyer, Rainer Raupach, Michael Lell, Bernhard Schmidt, and Marc Kachelrieß. Normalized metal artifact reduction (nmar) in computed tomography. Medical physics, 37(10):5482–5493, 2010.
[4] Esther Meyer, Rainer Raupach, Michael Lell, Bernhard Schmidt, and Marc Kachelrieß. Frequency split metal artifact reduction (fsmar) in computed tomography. Medical physics, 39(4):1904–1916, 2012.
[5] Haofu Liao, Wei-An Lin, Zhimin Huo, Levon Vogelsang, William J Sehnert, S Kevin Zhou, and Jiebo Luo. Generative mask pyramid network for ct/cbct metal artifact reduction with joint projection-sinogram correction. In International Conference on Medical Image Computing and Computer-Assisted Intervention, pages 77–85. Springer, 2019.
[6] Harshit Agrawal, Ari Hietanen, and Simo Särkkä. Metal artifact reduction in cone-beam extremity images using gated convolutions. In 2021 IEEE 18th International Symposium on Biomedical Imaging (ISBI), pages 1087–1090. IEEE, 2021.
[7] Lars Gjesteby, Hongming Shan, Qingsong Yang, Yan Xi, Bernhard Claus, Yannan Jin, Bruno De Man, and Ge Wang. Deep neural network for ct metal artifact reduction with a perceptual loss function. In In Proceedings of The Fifth International Conference on Image Formation in X-ray Computed Tomography, 2018.
[8] Xia Huang, Jian Wang, Fan Tang, Tao Zhong, and Yu Zhang. Metal artifact reduction on cervical ct images by deep residual learning. Biomedical engineering online, 17(1):1–15, 2018.
[9] Hong Wang, Yuexiang Li, Nanjun He, Kai Ma, Deyu Meng, and Yefeng Zheng. Dicdnet: Deep interpretable convolutional dictionary network for metal artifact reduction in ct images. IEEE Transactions on Medical Imaging, 41(4):869–880, 2021.
[10] Wei-An Lin, Haofu Liao, Cheng Peng, Xiaohang Sun, Jingdan Zhang, Jiebo Luo, Rama Chellappa, and Shaohua Kevin Zhou. Dudonet: Dual domain network for ct metal artifact reduction. In Proceedings of the IEEE/CVF Conference on Computer Vision and Pattern Recognition, pages 10512–10521, 2019.
[11] Tao Wang, Wenjun Xia, Yongqiang Huang, Huaiqiang Sun, Yan Liu, Hu Chen, Jiliu Zhou, and Yi Zhang. Dan-net: Dual-domain adaptive scaling non-local network for ct metal artifact reduction. Physics in Medicine & Biology, 66(15):155009, 2021.
[12] Lequan Yu, Zhicheng Zhang, Xiaomeng Li, and Lei Xing. Deep sinogram completion with image prior for metal artifact reduction in ct images. IEEE Transactions on Medical Imaging, 40(1):228–238, 2020.
[13] Sixiao Zheng, Jiachen Lu, Hengshuang Zhao, Xiatian Zhu, Zekun Luo, Yabiao Wang, Yanwei Fu, Jianfeng Feng, Tao Xiang, Philip HS Torr, et al. Rethinking semantic segmentation from a sequence-to-sequence perspective with transformers. In Proceedings of the IEEE/CVF conference on computer vision and pattern recognition, pages 6881–6890, 2021.
[14] Ze Liu, Yutong Lin, Yue Cao, Han Hu, Yixuan Wei, Zheng Zhang, Stephen Lin, and Baining Guo. Swin transformer: Hierarchical vision transformer using shifted windows. In Proceedings of the IEEE/CVF International Conference on Computer Vision, pages 10012–10022, 2021.
[15] Phillip Isola, Jun-Yan Zhu, Tinghui Zhou, and Alexei A Efros. Image-to-image translation with conditional adversarial networks. In Proceedings of the IEEE conference on computer vision and pattern recognition, pages 1125–1134, 2017.

## Slide 34: Thank you for your attention.
Any questions?